## BOM
RehabSense — Bill of Materials (complete two-leg system)
# | Component | Qty | Unit ₹ | Total ₹ | Vendor | Search link | Required? | Notes
1 | ESP32 Dev Board (38-pin DevKit V1) | 2 | 381 | 762 | Robocraze | https://robocraze.com/search?q=esp32+devkit | REQUIRED | One node per leg. 38-pin exposes GPIO32-39 (ADC1) needed for the FSR.
2 | GY-521 MPU6050 IMU module | 4 | 141 | 564 | Robocraze | https://robocraze.com/search?q=mpu6050 | REQUIRED | 2 per leg (thigh + shin). Buy the GY-521 breakout — it has the regulator and I2C pull-ups.
3 | SOUSHINE FSR402 force sensing resistor (14.7 mm, 20 g–10 kg) | 2 | 188 | 376 | robu.in / Amazon | https://robu.in/?s=FSR402 | OPTIONAL | CONFIRMED CORRECT. Direct FSR-402 equivalent. Saturates ~10 kg — irrelevant, the backend only needs loaded/unloaded for stance detection (threshold 0.15 normalised). SHORT TAIL: needs ~60-70 cm extension to reach the thigh node. Solder fast at low heat (the tail is flexible printed circuit and delaminates) and add strain relief.
4 | 10 kΩ resistor (pack of 10) | 1 | 13 | 13 | Robocraze | https://robocraze.com/search?q=10k+ohm+resistor | REQUIRED* | *Only if using the pressure sensor. Pull-down for the divider.
5 | 18650 Li-ion cell 2000–2600 mAh | 2 | 100 | 200 | Robocraze | https://robocraze.com/search?q=18650+battery | REQUIRED | One per node. 3.7 V nominal — must be boosted, see MT3608.
6 | TP4056 USB-C charge + protection module | 2 | 26 | 52 | Robocraze | https://robocraze.com/search?q=tp4056+usb+c | REQUIRED | Buy the version WITH protection circuit. Charges the 18650 safely.
7 | MT3608 step-up boost converter | 2 | 35 | 70 | Robocraze | https://robocraze.com/search?q=mt3608 | REQUIRED | Boost 3.7 V → 5 V into ESP32 VIN. Without this the board browns out. robu.in was out of stock.
8 | 18650 battery holder (single) | 2 | 17 | 34 | robu.in | https://robu.in/?s=18650+holder | REQUIRED | Sold as combo of 2 — one pack covers both nodes.
9 | SPST slide / toggle switch | 2 | 2 | 4 | robu.in | https://robu.in/?s=spst+slide+switch | REQUIRED | Power isolation per node.
10 | Perfboard 7×9 cm (pack of 2) | 1 | 59 | 59 | robu.in | https://robu.in/?s=perfboard+7x9 | REQUIRED | One board per node — the pack of 2 is exactly right.
11 | Male + female pin headers (40-pin strip) | 2 | 11 | 22 | robu.in | https://robu.in/?s=pin+header+40 | REQUIRED | Socket the ESP32 rather than soldering it down — you will want to remove it.
12 | 4-core shielded ribbon wire | 2 | 34 | 68 | robu.in | https://robu.in/?s=4+core+shielded+wire | REQUIRED | 1 m per leg: thigh ESP32 → shin IMU. Buy 2 m total.
13 | Plastic project enclosure | 2 | 40 | 80 | Robocraze | https://robocraze.com/search?q=plastic+enclosure+box | REQUIRED | Houses ESP32 + battery + boost, worn mid-thigh. One per leg.
14 | Heat-shrink tube assortment | 1 | 197 | 197 | robu.in | https://robu.in/?s=heat+shrink+tube | REQUIRED | 100 m roll is far more than needed; a small assorted pack is fine if cheaper.
15 | Zip ties (pack of 100) | 1 | 225 | 225 | robu.in | https://robu.in/?s=zip+ties | NICE | Cable management and strain relief.
ITEMS MISSING FROM YOUR LIST — add these
17 | Micro-USB / USB-C cable (data-capable) | 2 | 60 | 120 | Robocraze | https://robocraze.com/search?q=micro+usb+cable | MISSING | You CANNOT flash the ESP32 without this. Must be a DATA cable, not charge-only.
18 | Jumper wires M-M + M-F (40 each) | 2 | 40 | 80 | robu.in | https://robu.in/?s=jumper+wires | MISSING | Essential for breadboarding before you solder anything.
19 | Breadboard 830-point | 2 | 70 | 140 | robu.in | https://robu.in/?s=breadboard+830 | MISSING | Prove the I2C wiring works before committing to perfboard.
20 | Elastic velcro strap 5 cm × 1 m | 4 | 50 | 200 | robu.in | https://robu.in/?s=velcro+strap+elastic | MISSING | Replaces the ₹382 neoprene bands you couldn't find — cheaper and better for sensor mounting.
21 | JST-XH 4-pin connector set | 1 | 60 | 60 | robu.in | https://robu.in/?s=jst+xh+4+pin | NICE | Makes the shin sensor cable detachable. A soldered run will fail at the knee.
99 | Thin flexible wire for FSR tail extension (2-core, ~1.5 m) | 1 | 30 | 30 | robu.in | https://robu.in/?s=silicone+wire+26awg | REQUIRED* | *If using the short-tail FSR. Runs heel → thigh node. Thin and flexible so it survives walking.
TOTAL

## Wiring
RehabSense — wiring, per leg node (build this twice)
From | To | Why it matters
18650 (+) via switch | TP4056 B+ | Charge and protection sit closest to the cell.
TP4056 OUT+ | MT3608 IN+ | Boost the 3.7 V cell up before the ESP32.
MT3608 OUT+ (set to 5.0 V) | ESP32 VIN | SET THE OUTPUT TO 5 V WITH A MULTIMETER BEFORE CONNECTING. The trimpot ships arbitrary.
ESP32 3V3 | Both MPU6050 VCC | GY-521 has an onboard regulator; 3.3 V is correct.
ESP32 GPIO21 (SDA) | Both MPU6050 SDA | Shared I2C bus.
ESP32 GPIO22 (SCL) | Both MPU6050 SCL | Shared I2C bus.
Thigh MPU6050 AD0 | GND | Sets I2C address 0x68.
Shin MPU6050 AD0 | 3.3 V | Sets I2C address 0x69 — this is what allows two IMUs on one bus.
Pressure sensor leg 1 | ESP32 3V3 | Top of the voltage divider.
Pressure sensor leg 2 | GPIO34 + 10 kΩ to GND | MUST be an ADC1 pin (32-39). ADC2 is dead while Wi-Fi is on.
All grounds | Common GND | Battery, boost, ESP32 and both IMUs share one ground.

## Firmware contract
What the firmware must send — the backend rejects anything else
Field | Requirement
Transport | Wi-Fi WebSocket to ws://<server>:8000/ws/ingest/{session_id}?leg=LEFT or RIGHT
Handshake | Send {"type":"hello","protocol_version":1,"device_id":"...","leg":"LEFT","firmware_version":"...","sensors":["thigh_imu","shin_imu","fsr"]}
Server replies | hello_ack with session_id, server_time, sample_rate_hz (100.0)
Sample rate | 100 Hz per leg
Data packet | {"type":"data","leg":"LEFT","samples":[ ... ]}
Each sample | {"ts": seconds_float, "seq": int, "thigh":{ax,ay,az,gx,gy,gz}, "shin":{...}, "fsr": 0.0-1.0 optional}
Accelerometer units | g (NOT m/s²)
Gyroscope units | deg/s (NOT rad/s)
seq | Per-SAMPLE counter, monotonically increasing. Used to detect dropped packets.
Calibration | Hold both limbs still for the first 2.5 s — the server measures gyro bias then.
simulated flag | Real hardware must OMIT it (or send false). Only the simulator sets it true.
Batching | Send ~10 samples per packet at ~10 packets/s. Verified stable to 4000 samples/s.

## Risks
Things that will cost you a day if missed
Risk | What to do
College Wi-Fi blocks devices | HIGHEST RISK. Most campus networks isolate clients, so the ESP32 cannot reach your laptop. Test with a phone hotspot BEFORE demo day. This kills the demo regardless of the electronics.
MT3608 output not set | It ships at an arbitrary voltage. Turn the trimpot and confirm 5.0 V on a multimeter BEFORE wiring it to the ESP32, or you can destroy the board.
FSR on ADC2 | GPIO 0/2/4/12-15/25-27 are ADC2 and return garbage while Wi-Fi is active. Use GPIO 32-39 only.
Both IMUs on 0x68 | If you forget the AD0 pin, only one IMU responds and knee angle is impossible. Run an I2C scanner first — you must see 0x68 AND 0x69.
Charge-only USB cable | A cable without data lines will not flash the board and gives no error that explains why.
Loose sensor mounting | A sensor that slides on the skin adds error the filter cannot remove. Mount rigidly to the strap, not to loose fabric.
No spare MPU6050 | They are cheap and easy to kill with a wiring mistake. Buy one or two extra.